Take Home Instructions

Getting started at home

These directions will be used every time the oximeter unit is unplugged (or forced shutdown). The oximeter will already be configured to your child’s pre-determined threshold parameters for heart rate and oxygen saturation.

Step 1:

*The oximeter is already set-up by staff and is ready for use.

- If not already plugged in by staff, plug the Patient Cable into the front of the oximeter (Figure 1, Arrow “A”)

- Plug the oximeter’s thick black electrical cord into a standard wall outlet (Note: the oximeter has a 2-hour battery life)

- POWER ON the oximeter by pressing the power button until you hear a “beep” (Figure 1, Arrow “B”)

- The green light at the top will remain lit to indicate that the device is on (Figure1, Arrow “C”)

A

A

C

C

B

B

Figure 1

Step 2: Configure the oximeter to your home WIFI network

| - Tap the icon on the top of the |  |
| --- | --- |
| - Press the “i” icon |  |
| - Select your home network from the list |  |
| - If applicable, enter network password |  |
| - Press “Save” |  |

Troubleshooting

| Wrong Password |  |
| --- | --- |

Step 3: Now it is ready to wear

- Attach the probe to the top of your child’s foot or to the hand as described

- Plug the probe into the other end of the patient cable (the grey arrow on the probe and the white arrow on the patient cable will align)

Sensor Placement

The sensor

- The sensor is the device at the end of the white patient cable that uses a red light and a detector to measure oxygen saturation

Where to place sensor

- The sensor should be placed on the foot or hand as described during your training

- The sensor emitter (red light) and the detector should be properly aligned with the detector directly opposite the red light. There are designations on the adhesive tape to help you with the alignment.

When to change sensor location

- The sensor should be rotated from one site to another site twice a day if wearing continuously

When and how to use replacement tapes

- If the adhesive tape does not stick any more, you can carefully remove the adhesive tape from the sensor and replace it with a new piece of tape

A few things to note:

- The oximeter only has about a two-hour battery life

- Each time the probe is disconnected the oximeter will alarm until it is reconnected

- Each time the oximeter enters a new WiFi network location you will need to repeat Step 2 to configure to network

- To power off you need to hold down the power button (Figure 1, Arrow “B”) for 10 seconds until you hear the “beep”

Troubleshooting

| INSTRUMENT DOES NOT POWER ON | - Connect the oximeters thick black power cord to a wall outlet and to the back of the oximeter. - Make sure that the light on top of the oximeter is green. |
| --- | --- |
| “SENSOR OFF” MESSAGE (“Sensor off patient” or “Searching for probe”) | - Reapply the probe to the baby and reconnect the sensor/probe to the patient cable, which is plugged into the front of the oximeter. - If message does not go away by repositioning, replace the current probe with a new one. - Please note that the “Sensor Off” message will display temporarily each time the probe is placed on the infant while the machine is searching for patient’s pulse (about 10 seconds after placing probe) |
| “PULSE SEARCH” message | - If the “PULSE SEARCH” message displays for more than 30 seconds, disconnect and reconnect sensor/probe cord. |
| DISPLAYING DASHES (numeric values missing) | - Assess baby and check the placement of the sensor. Re-apply sensor or move to another site. - Replace sensor/probe. |
| UNEXPECTED SpO2 READING | - Assess the baby. - Reposition sensor/probe to site with a strong signal quality. |
| WIFI Symbol is “Blue” | - Contact your home care company as this needs to be solved by them. |

Frequent Questions

When will I receive results of the recorded home oximetry?

Data is downloaded from the system every [Monday and Thursday/Tuesday and Friday]. You will be contacted by [Site Personnel Name] or other provider with the results of your data. The results will either be a maintain, increase, or decrease of your child’s current oxygen flow rate.

Do I need to be on WIFI for the data to be recorded and analyzed?

Yes, the oximeter needs to be connected to WIFI for data to be recorded by the system. If you do not have WIFI or are going to be spending a long period of time away from WIFI contact the team for instructions.

If I am having trouble with the oximeter of have a question who do I contact?

Contact [Contact Name], his/her contact information is listed below.

Contact Info

Please don’t hesitate to contact us with questions. We would be more than happy to help!

[Insert your site’s contact information here]